BD LSRFortessaTM SORP

**Description**

The BD LSRFortessaTM SORP is a benchtop digital flow cytometer. Equipped with 5 spatially separated lasers, 18 fluorescent detectors and 2 light scatter detectors for fast multicolour analysis with accurate data acquisition rate at up to 40,000 events per second (with beads). The BD LSRFortessaTM has a BDTM High Throughput Sampler (HTS) for increased lab productivity by acquiring samples from a 96- or 384-well plates.

**Applications**

Multicolour flow cytometry analysis.

**Excitation Optics**

Laser power

355 nm: 20 mW

405 nm: 50 mW

488 nm: 100 mW

561 nm: 50 mW


640 nm: 40 mW

**Emission Optics**


Forward Scatter Detector: photodiode with 488/10 bandpass filter for the 488-nm laser.

Side Scatter Detector: photomultipler with a 488/10 bandpass filter for the 488-nm laser.

Fluorescence Detectors and Filters:

****

Available filters



**Fluidics**

Sample flow rates

Front key panel provides three modes: RUN, STANDBY, and PRIME

Continuously adjustable flow rates, plus three preset flow rates:

LO: 12 µL/min

MED: 35 µL/min

HI: 60 µL/min

**Performance**

Fluorescence Sensitivity

* FITC : 80 molecules of equivalent soluble fluorochrome (MESF-FITC)
* PE: 30 molecules of equivalent soluble fluorochrome (MESF-PE)
* PE-Cy™5: 10 molecules of equivalent soluble fluorochrome (MESF-PE-Cy5)
* APC: 70 molecules of equivalent soluble fluorochrome (MESF-APC)

FITC and PE measurements performed using SPHERO™ Rainbow Calibration Particles (RCP-30-5A) PE-Cy5 and APC measurements performed using SPHERO Ultra Rainbow Calibration Particles (URCP-38-2K)

Fluorescence Resolution

Coefficient of variation (CV) PI: Area, <3.0%, full G0/G1 peak for PI-stained chicken erythrocyte nuclei (CEN)

Fluorescence Linearity

Doublet/singlet ratio CEN stained with PI: 1.95–2.05 (488-nm laser)

Forward and Side Scatter Sensitivity

Sensitivity enables separation of fixed platelets from noise, identification of bacteria, and detection of 0.5-µm beads.

Forward and Side Scatter Resolution

Scatter performance is optimized for resolving lymphocytes, monocytes, and granulocytes.

**Signal Processing**

Converter

10-MHz analog-to-digital converter.

Workstation Resolution 262,144 channels

Pulse Processing: height, area, and width measurements available for any parameter. Ratio measurements are also available.

Time can be correlated to any parameter for kinetic experiments or other applications.

**BD™ High Throughput Sampler (HTS)**

is available to increase your lab productivity by acquiring samples from a 96- or 384-well microtiter plate.

HTS throughput

Acquisition: less than 15 minutes per microtiter plate in high throughput mode using a 2-second acquisition,

less than 44 minutes in standard mode using a 10-second acquisition

Carryover: <0.5% HT mode

<0.75% STD mode

Link to the manufacturer [BD Biosciences](http://www.bdbiosciences.com/us/instruments/research/cell-analyzers/bd-lsrfortessa/m/751752/overview)

Recommended tool for multicolour experiment design: [SpectrumViewer](http://www.bdbiosciences.com/us/s/spectrumviewer)